PROPULSE MAINTENCE

Propulse Weekly Cleaning Record

WGPN

|  |  |  |  |
| --- | --- | --- | --- |
| **Date/Month** | **Cleaned** | **Name** | **Signature** |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |
|  |  |  |  |



**Propulse Cleaning Guidelines**

Do not attempt to clean the Jet Tip. Use one Propulse Disposable 'OneUse' Jet Tip / Propulse Single Use QrX™ per patient treatment and

discard to clinical waste after use as this reduces the risk of cross

infection between patients.

External cleaning should be done by hand, wiping with a damp cloth

only. Apply liquids to cloth not the unit. Do not immerse the unit in

water. Mild detergents and disinfectants may be used externally.

Internal cleaning of the Propulse II / III / NG Ear Irrigator pump and

waterway must be carried out PRIOR to use.

* Place one Propulse Cleaning Tablet into the reservoir, fill with warm water to the 500ml mark on the reservoir and wait for thetablet to completely dissolve.
* Run the Propulse Ear Irrigator for a few seconds to fill the pump
* and hose.
* Leave to stand for 10 minutes.
* Empty the reservoir of the solution, then fill the reservoir with cool

boiled water or well run cold tap water and operate the Propulse to

flush through the whole system to ensure no cleaning solution

remains.

* This cleaning procedure should be done weekly.
* DO NOT leave the solution in the unit for longer than stated.
* Cleaning Tablets for the Propulse Electronic Ear Irrigator - the

importance of using the correct strength cleaning solution cannot

be overstated. A solution that is too strong will in time damage

the Propulse and one that is too weak will fail to provide the correct

level of cleaning and decontamination. To this end the manufacturer

has introduced a simple to use, effective specific strength

* CHLOR-CLEAN™ decontamination tablet. Propulse Cleaning Tablets

 are available from your normal Propulse supplier or direct from AMA suppliers.

Please Note: Damage caused to the Propulse Ear Irrigator by not using the recommended cleaning agent will invalidate our warranty. Other instruments used with the Propulse should be cleaned according to local guidelines.